******

2nd Floor Malong Building, Capitol Compound

Lingayen, Pangasinan

Tel. No. (075) 632-7840

Website@biddingandawards.pangasinan gov.ph, E-mail address: [bacpangasinan@yahoo.com](mailto:bacpangasinan@yahoo.com)

**INVITATION TO BID**

**for**

**Supply and Delivery of Ten (10) units X-Ray Portable at Provincial Governor’s Office, Lingayen, Pangasinan (for use of various hospitals)**

Project Identification No.: PANG-2024-01-0005-G

1. The **Provincial Government of Pangasinan**, through the **LBP Loan (PR#2024-01-0019)** intends to apply the sum of **One Hundred Fifty Million Pesos (P150,000,000.00)** being the Approved Budget for the Contract (ABC) to payments under the contract for **Supply and Delivery of Ten (10) units X-Ray Portable at Provincial Governor’s Office, Lingayen, Pangasinan (for use of hospitals).**

**Specifications:**

Collapsible Column

Column Height 127cm during transportation, Focus Height 680 2025mm consisting of:

**High Frequency Inverter High Voltage Generator** ­

• Maximum Inverter Frequency : 60kHz

­• Tube Voltage Range : 40-133kV at 1 kV Step

• Maximum Tube Current : 400mA at 20msec

• mAs Range : 0.32-320mAs at 12.5% Steps

• Maximum Output : 32 kW ­

• Anatomical Programs: 441 (432 for FPD Radiography and 9 for General Radiography)

Status Indicator: 4 Colors

**Main Mobile Cart with Telescopic-Arm Support Column**

• Power Assisted Optimization System "PAOS" "

• Inch Mover" Movement Control at Collimator Panel

• Arm Rotation : +/- 270 degrees

• Arm Horizontal Travel : 120 cm • Narrow Width : 56 cm

**Battery**

**X-ray Tube Unit** ­

• Anode Heat Capacity : 300kHU ­

• Focal Spots : 0.7 / 1.3 mm Dual-Focus

**Multi-Leaf LED Collimator Compartment for Canon FPD Illuminated Hand Switch**

**Display: 17inch Touch-screen LCD**

**Key-less Entry**

**Digital Radiography (FPD) consisting of:**

**Flat Panel Detector: Wireless and Cassette-Size**

• Wireless Standard : IEEE 802.11n ­

• Image Size : Auto-Sizing up to 17" x 14" ( 35 x 43cm )

• Scintillator : Caesium Iodide ( CsI )

• "Cassette-Size" Sensor Dimension : 384 x 460 x 15 mm

• Preview Image Time : 2 seconds after exposure

• Acquisition Depth : 16-bits

• Battery Performance : 140 Images @100sec-cycle

• Sensors Identification : Up to 4-Active, 10-Registered ­ Sensor

• Exposure Time : Up to 3 seconds

• Sensor Weight : 3.2 kg / 7.3 lbs ( Battery Included )

• Dose Area Product meter

• Infrared Remote controller

• Decoration Labe

**CXDI Control Software ( NE )**

Includes DICOM Print / Store / MWM / MPPS

Generator Communication Module CXDI GCOM

X-ray Interface Unit Infrared Data Communication Unit

Battery Charger

**Accessories Included: FOC**

1. Lead Gown 1 pc

2. Thyroid Shield 1 pc

3. Chest Thickness Caliper 1 pc

4. Gonadal Shield 1 set

Bids received in excess of the ABC shall be automatically rejected at bid opening.

1. The **Provincial Government of Pangasinan** now invites bids for **Supply and Delivery of Hospital Equipment.** Delivery of the Goods is required Seven (7) Calendar Days. Bidders should have completed from the date of submission and receipt of bids, a contract similar to the project. The description of an eligible bidder is contained in the Bidding Documents, particularly, in Section II. Instruction to Bidders.
2. Bidding will be conducted through open competitive bidding procedures using a non-discretionary “pass/fail” criterion as specified in the 2016 Revised Implementing Rules and Regulation (IRR) of Republic Act (RA) 9184), otherwise known as the “Government Procurement Reform Act”

Bidding is restricted to Filipino citizen/sole proprietorships, partnerships, or organizations with at least sixty percent (60% interest or outstanding capital stock belonging to citizens of the Philippines, and to citizens or organizations of a country the laws or regulations of which grant similar rights or privileges to Filipino citizens, pursuant to RA 5183.

1. Interested bidders may obtain further information from **Provincial Government of Pangasinan** and inspect the Bidding Documents at the address given below during Bidders **January 6, 2024 – January 25, 2024; 8:00 am to 5:00pm and January 26, 2024; 8:00 am to 10:00am**.
2. A complete set of Bidding Documents may be acquired by interested Bidders during **January 6, 2024 – January 25, 2024; 8:00 am to 5:00pm and January 26, 2024; 8:00 am to 10:00am at the BAC Office, 2nd Floor Malong Building, Lingayen, Pangasinan** and upon payment of the applicable fee for the Bidding Documents, pursuant to the latest Guidelines issued by the GPPB, in the amount of **Fifty Thousand Pesos (P50,000.00)**.

It may also be downloaded free of charge from the website of the Philippine Government Electronic Procurement System (PhilGEPS) and the website of the Procuring Entity, provided that Bidders shall pay the applicable fee for the Bidding Documents not later than the submission of their bids.

1. The **Provincial Government of Pangasinan** will hold a **Pre-Bid Conference** on **January 15, 2024; 2:00 pm** at **Conference Room, 2nd Floor Malong Building, Capitol Compound, Lingayen, Pangasinan**, which shall be open to prospective bidders.
2. Bids must be duly received by the BAC Secretariat at the address below on or before **January 26, 2024; 10:00am**. All bids must be accompanied by a bid security in any of the acceptable forms and in the amount stated in ITB Clause 18

Bid opening shall be on **January 26, 2024; 10:00am** at **Conference Room, 2nd Floor, Malong Building, capitol Compound, Lingayen, Pangasinan**. Bids will be opened in the presence of the bidders’ representatives who choose to attend at the address below. Late bids shall not be accepted.

1. The **Provincial Government of Pangasinan** reserves the right to reject any and all bids, declare a failure of bidding, or not award the contract at any time prior to contract award in accordance with Section 41 of RA 9184 and its IRR, without thereby incurring any liability to the affected bidder or bidders.
2. For further information, please refer to:

**MELICIO F. PATAGUE II**

Provincial Administrator

BAC Chairman

Provincial Administrator

Capitol Building, Capitol Compound

Lingayen, Pangasinan

**MARLON C. OPERAÑA**

Provincial Accountant

BAC Technical Working Group

Provincial Accountant

Finance Building, Capitol Compound

Lingayen, Pangasinan

**ROWENA V. IGNACIO**

OIC – Provincial Planning & Dev’t Office

BAC Secretariat

BAC Office

2nd Floor Malong Building, Capitol Compound

Lingayen, Pangasinan

075 6327840

\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

**MELICIO F. PATAGUE II**

Provincial Administrator

BAC Chairman